Receiver Model: TDEMI_v5.11_1608003

Protocol Number: A75

Serial Number: 10

Operation State:

Notes:

Traces:

| Trace | Start | Step | Stop | IF Bandwidth | Detector | Dwell time | Attenuation | Transducer | LISN | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Scan 1 | 9.0 kHz | 100.0 Hz | 150.0 kHz | 200 Hz | QP | 1.0 s | Auto | TDF1, TDF2, TDF3 | NSLK8127-N |  |
| Scan 2 | 150.0 kHz | 5.0 kHz | 30.0 MHz | 9 kHz | QP | 1.0 s | Auto | TDF1, TDF2, TDF3 | NSLK8127-N |  |
| Scan 3 | 150.0 kHz | 5.0 kHz | 30.0 MHz | 9 kHz | Average | 100.0 ms | Auto | TDF1, TDF2, TDF3 | NSLK8127-N |  |
| Scan 4 | 9.0 kHz | 100.0 Hz | 150.0 kHz | 200 Hz | QP | 1.0 s | Auto | TDF1, TDF2, TDF4 | NSLK8127-L1 |  |
| Scan 5 | 150.0 kHz | 5.0 kHz | 30.0 MHz | 9 kHz | QP | 1.0 s | Auto | TDF1, TDF2, TDF4 | NSLK8127-L1 |  |
| Scan 6 | 150.0 kHz | 5.0 kHz | 30.0 MHz | 9 kHz | Average | 100.0 ms | Auto | TDF1, TDF2, TDF4 | NSLK8127-L1 |  |

Transducers:

TDF1: Attenuator-Tab 10dB - Serial No. 170201

TDF2: BNC Cable DC - 1GHz  -  1.5m

TDF3: Schwarzbeck NSLK 8127-RC05021 Phase-N (9kHz-30MHz)

TDF4: Schwarzbeck NSLK 8127-RC05021 Phase-L (9kHz-30MHz)

Limit lines:

EN55015,QP: EN55015 QP, 9kHz-30MHz

EN55015,AV: EN55015 Average,150kHz-30MHz

Scan1: 9.0 kHz, 100.0 Hz, 150.0 kHz; IF:200 Hz, 1.0 s QP, Att AutodB, LISN-NSLK8127-N

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 9.615 kHz | 58.58 | 110.00 | 51.42 | 10.80 | EN55015,QP |  |
| 62.295 kHz | 37.61 | 88.00 | 50.39 | 10.00 | EN55015,QP |  |
| 76.983 kHz | 44.10 | 86.07 | 41.98 | 10.00 | EN55015,QP |  |
| 121.203 kHz | 25.60 | 81.94 | 56.34 | 10.00 | EN55015,QP |  |
| 150.000 kHz | 21.82 | 66.00 | 44.18 | 10.00 | EN55015,QP |  |

Scan2: 150.0 kHz, 5.0 kHz, 30.0 MHz; IF:9 kHz, 1.0 s QP, Att AutodB, LISN-NSLK8127-N

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 1.376 MHz | 50.64 | 56.00 | 5.36 | 10.02 | EN55015,QP |  |
| 6.420 MHz | 45.76 | 60.00 | 14.24 | 10.09 | EN55015,QP |  |
| 12.159 MHz | 38.42 | 60.00 | 21.58 | 10.22 | EN55015,QP |  |
| 18.164 MHz | 31.42 | 60.00 | 28.58 | 10.24 | EN55015,QP |  |
| 24.056 MHz | 27.10 | 60.00 | 32.90 | 10.37 | EN55015,QP |  |

Scan3: 150.0 kHz, 5.0 kHz, 30.0 MHz; IF:9 kHz, 100.0 ms Average, Att AutodB, LISN-NSLK8127-N

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 381.328 kHz | 32.85 | 48.25 | 15.40 | 10.00 | EN55015,AV |  |
| 6.194 MHz | 25.15 | 50.00 | 24.85 | 10.09 | EN55015,AV |  |
| 12.159 MHz | 20.76 | 50.00 | 29.24 | 10.22 | EN55015,AV |  |
| 18.203 MHz | 15.96 | 50.00 | 34.04 | 10.24 | EN55015,AV |  |
| 24.247 MHz | 12.69 | 50.00 | 37.31 | 10.37 | EN55015,AV |  |

Scan4: 9.0 kHz, 100.0 Hz, 150.0 kHz; IF:200 Hz, 1.0 s QP, Att AutodB, LISN-NSLK8127-L1

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 9.615 kHz | 58.86 | 110.00 | 51.14 | 10.85 | EN55015,QP |  |
| 62.372 kHz | 37.88 | 87.99 | 50.11 | 10.00 | EN55015,QP |  |
| 76.983 kHz | 44.04 | 86.07 | 42.03 | 10.00 | EN55015,QP |  |
| 121.357 kHz | 26.41 | 81.93 | 55.52 | 10.00 | EN55015,QP |  |
| 150.000 kHz | 22.17 | 66.00 | 43.83 | 10.00 | EN55015,QP |  |

Scan5: 150.0 kHz, 5.0 kHz, 30.0 MHz; IF:9 kHz, 1.0 s QP, Att AutodB, LISN-NSLK8127-L1

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 563.438 kHz | 51.24 | 56.00 | 4.76 | 10.00 | EN55015,QP |  |
| 6.878 MHz | 45.22 | 60.00 | 14.78 | 10.10 | EN55015,QP |  |
| 12.150 MHz | 38.36 | 60.00 | 21.64 | 10.23 | EN55015,QP |  |
| 18.149 MHz | 31.18 | 60.00 | 28.82 | 10.25 | EN55015,QP |  |
| 24.031 MHz | 27.39 | 60.00 | 32.61 | 10.39 | EN55015,QP |  |

Scan6: 150.0 kHz, 5.0 kHz, 30.0 MHz; IF:9 kHz, 100.0 ms Average, Att AutodB, LISN-NSLK8127-L1

| f | Mag [dBµV] | Limit | Diff | Trans | Name | Comment |
| --- | --- | --- | --- | --- | --- | --- |
| 376.406 kHz | 35.35 | 48.36 | 13.01 | 10.00 | EN55015,AV |  |
| 6.199 MHz | 29.97 | 50.00 | 20.03 | 10.09 | EN55015,AV |  |
| 12.154 MHz | 22.77 | 50.00 | 27.23 | 10.23 | EN55015,AV |  |
| 18.297 MHz | 18.86 | 50.00 | 31.14 | 10.26 | EN55015,AV |  |
| 24.159 MHz | 16.23 | 50.00 | 33.77 | 10.39 | EN55015,AV |  |